## Slide 1: Two reactions
STOP
8 faults stop the session
Enter Error
Pump stops
Compressor turns off
Orange error chip
Acknowledge returns to Ready
MESSAGE
4 warnings leave it running
State stays the same
Pump keeps running
Compressor stays on
Yellow warning banner
Banner clears when the cause clears
Fridge defect and battery low have no live reading yet. Leak detection is off in config.yaml.
Notes: Stop faults enter Error, stop the pump, and turn the compressor off. Acknowledge returns to Ready. Message faults only show a yellow banner. If several stop faults are active, only the highest-priority one is latched. Source: src/fault_catalog.py, src/safety_rules.py, src/main.py.

## Slide 2: Stop faults, highest priority
Shared reaction: enter Error, stop the pump, and turn the compressor off. Acknowledge returns to Ready.
No. | Message | When it trips | Acknowledge
1 | Sensor read failure | A sensor read throws. The chip shows the exception text. | Allowed immediately. The next failed read trips again.
2 | Level sensor failure detected | Cooling or pumping, and a level switch is not OK. | Both level switches read OK.
3 | Cartridge removed during operation | Cooling or pumping, and the cartridge is not seated. | Cartridge switch reads ON.
4 | Fridge defect | A fridge-defect flag. The live tick leaves that flag empty. | The defect flag clears.
Notes: Priority order when several stop faults are active: I/O, level, cartridge, fridge, leak, CSF low, heat exchanger, cooling ineffective. Only the first one is latched.

## Slide 3: Stop faults, continued
Shared reaction: enter Error, stop the pump, and turn the compressor off. Acknowledge returns to Ready.
No. | Message | When it trips | Acknowledge
5 | Leak detected | Leak line stays low for 0.5 s. Detection is off in config. | The leak sensor reads dry.
6 | CSF low temp | Cooling or pumping, and Body Temp is below 28 °C. | Body Temp is at least 28 °C.
7 | Heat exchanger too cold | Cooling or pumping, and Plate 1 is below -10 °C. | Plate 1 is at least -10 °C.
8 | Cooling ineffective | Pump and compressor on, Body Temp falls under 0.2 °C in 10 min. | Next tick. Stopping the pump clears the rule.
Notes: Thresholds from config.yaml: Body Temp 28 C, Plate 1 -10 C, leak debounce 0.5 s, cooling ineffective 600 s and 0.2 C. leak_detection_enabled is false.

## Slide 4: Warnings
Shared reaction: yellow banner only. The session keeps running. No acknowledge.
Message | When it trips | What you see
Battery low | Reported charge below 20%. The live tick leaves battery percent empty. | Banner. Session continues.
USB stick not present | USB logging is on and no SPINELOGS stick is mounted. | Banner. Logged once per session.
SD card storage below 1 GB | The local logs disk has under 1 GB free. | Banner. Clears once 1 GB is free.
USB stick storage below 1 GB | A mounted stick has under 1 GB free. | Banner. Clears once 1 GB is free.
Notes: USB stick not present is written once per session. A missing stick does not also raise the USB storage warning. Simulation skips USB presence checks.