## Faults
Fault messages and reactions
Stop faults | 8 | Warnings | 4
Messages are the strings in src/fault_catalog.py. Trip conditions and reactions follow src/safety_rules.py and src/main.py. Blue values on the Limits sheet are the config.yaml thresholds; sentences that use them update from that sheet. Stop rows are in latch priority order. If several stop faults are active, only the first one is shown.
No. | Code | Message | Class | When it trips | Reaction | Shared reaction
1 | IO_READ_FAILURE | Sensor read failure | Stop | A sensor read throws. The chip shows the exception text. | Acknowledge is allowed immediately. The next failed read trips again. | Enter Error. Stop the pump and turn the compressor off. Acknowledge returns to Ready.
2 | LEVEL_SENSOR | Level sensor failure detected | Stop | Cooling or pumping, and a level switch is not OK. | Acknowledge stays off until both level switches read OK. | Enter Error. Stop the pump and turn the compressor off. Acknowledge returns to Ready.
3 | CARTRIDGE_REMOVED | Cartridge removed during operation | Stop | Cooling or pumping, and the cartridge is not seated. | Acknowledge stays off until the cartridge switch reads ON. | Enter Error. Stop the pump and turn the compressor off. Acknowledge returns to Ready.
4 | FRIDGE_DEFECT | Fridge defect | Stop | A fridge-defect flag. The live tick leaves that flag empty. | Acknowledge stays off until the defect flag clears. | Enter Error. Stop the pump and turn the compressor off. Acknowledge returns to Ready.
5 | LEAK_DETECTED | Leak detected | Stop | Leak line stays low for 0.5 s. Leak detection is off. | Acknowledge stays off until the leak sensor reads dry. | Enter Error. Stop the pump and turn the compressor off. Acknowledge returns to Ready.
6 | CSF_LOW_TEMP | CSF low temp | Stop | Cooling or pumping, and Body Temp is below 28 °C. | Acknowledge stays off until Body Temp is at or above 28 °C. | Enter Error. Stop the pump and turn the compressor off. Acknowledge returns to Ready.
7 | HEAT_EX_TOO_COLD | Heat exchanger too cold | Stop | Cooling or pumping, and Plate 1 is below -10 °C. | Acknowledge stays off until Plate 1 is at or above -10 °C. | Enter Error. Stop the pump and turn the compressor off. Acknowledge returns to Ready.
8 | COOLING_INEFFECTIVE | Cooling ineffective | Stop | Pump and compressor on, and Body Temp falls less than 0.2 °C within 600 s. | Acknowledge is available on the next tick. Stopping the pump clears the rule. | Enter Error. Stop the pump and turn the compressor off. Acknowledge returns to Ready.
9 | BATTERY_LOW | Battery low | Message | Reported charge is below 20%. The live tick leaves battery percent empty. | Yellow banner. The session continues. | Yellow banner only. The session keeps running. No acknowledge.
10 | USB_NOT_PRESENT | USB stick not present | Message | USB logging is on and no SPINELOGS stick is mounted. | Yellow banner. Logged once per session. | Yellow banner only. The session keeps running. No acknowledge.
11 | SD_STORAGE_LOW | SD card storage below 1 GB | Message | Free space on the local logs disk is below 1 GB. | Yellow banner. Clears once free space is back above 1 GB. | Yellow banner only. The session keeps running. No acknowledge.
12 | USB_STORAGE_LOW | USB stick storage below 1 GB | Message | A mounted stick has less than 1 GB free. | Yellow banner. Clears once free space is back above 1 GB. | Yellow banner only. The session keeps running. No acknowledge.
Shared reaction follows Class. Fridge defect and battery low have no live reading yet. Leak detection is off in config.yaml.

## Limits
Limits used by the fault list
Yellow cells with blue text are the config.yaml values as of this list. The Faults sheet reads them. The on-screen message strings stay on the Faults sheet because those are fixed in the fault catalog.
Setting | Value | Unit | Source
CSF sensor | Body Temp |  | config.yaml alarms.csf_label
CSF low temp | 28 | °C | config.yaml alarms.csf_low_temp_c
Heat exchanger sensor | Plate 1 |  | config.yaml alarms.heat_ex_label
Heat exchanger minimum | -10 | °C | config.yaml alarms.heat_ex_min_c
Cooling ineffective timeout | 600 | s | config.yaml alarms.cooling_ineffective_timeout_s
Cooling ineffective drop | 0.2 | °C | config.yaml alarms.cooling_ineffective_csf_delta_c
Battery warning | 20 | % | config.yaml alarms.battery_low_pct
Storage minimum | 1 | GB | config.yaml alarms.storage_min_free_gb
Leak debounce | 0.5 | s | config.yaml alarms.leak_debounce_s
Leak detection | Off |  | config.yaml alarms.leak_detection_enabled is false